## summary
regression | X | Y | n | r | slope | intercept | BA_bias | n_outliers
PassingBablok | TAT207 | TAT1 | 81 | 0.98534 | 0.75904 | 1.1386 | -1.3494 | 6

## sample_metrics
SampleID | TAT207 | TAT1 | X | Y | diff_y_minus_x | abs_diff_yx | mean_xy | rel_diff_yx_pct | bias_pct_vs_x | ratio_y_over_x | pred_y | residual | abs_residual | z_MAD | abs_z_MAD | outlier_level | outlier_basis | outlier_note | regression
| 13.6 | 11 | TAT207 | TAT1 | -2.6 | 2.6 | 12.3 | 21.1382 | -19.1176 | 0.80882 | 11.4614 | -0.46145 | 0.46145 | -0.43563 | 0.43563 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 7.7 | 5.8 | TAT207 | TAT1 | -1.9 | 1.9 | 6.75 | 28.1481 | -24.6753 | 0.75325 | 6.9831 | -1.1831 | 1.1831 | -1.1169 | 1.1169 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 12.2 | 9.1 | TAT207 | TAT1 | -3.1 | 3.1 | 10.65 | 29.108 | -25.4098 | 0.7459 | 10.3988 | -1.2988 | 1.2988 | -1.2261 | 1.2261 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 13 | 9.7 | TAT207 | TAT1 | -3.3 | 3.3 | 11.35 | 29.0749 | -25.3846 | 0.74615 | 11.006 | -1.306 | 1.306 | -1.233 | 1.233 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 6.9 | 7.2 | TAT207 | TAT1 | 0.3 | 0.3 | 7.05 | 4.2553 | 4.3478 | 1.0435 | 6.3759 | 0.8241 | 0.8241 | 0.778 | 0.778 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 21.5 | 20.1 | TAT207 | TAT1 | -1.4 | 1.4 | 20.8 | 6.7308 | -6.5116 | 0.93488 | 17.4578 | 2.6422 | 2.6422 | 2.4944 | 2.4944 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 4.7 | 4.4 | TAT207 | TAT1 | -0.3 | 0.3 | 4.55 | 6.5934 | -6.383 | 0.93617 | 4.706 | -0.30602 | 0.30602 | -0.2889 | 0.2889 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 11.1 | 9.2 | TAT207 | TAT1 | -1.9 | 1.9 | 10.15 | 18.7192 | -17.1171 | 0.82883 | 9.5639 | -0.36386 | 0.36386 | -0.3435 | 0.3435 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 24.3 | 23.1 | TAT207 | TAT1 | -1.2 | 1.2 | 23.7 | 5.0633 | -4.9383 | 0.95062 | 19.5831 | 3.5169 | 3.5169 | 3.3201 | 3.3201 | mild_candidate | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 30 | 16 | TAT207 | TAT1 | -14 | 14 | 23 | 60.8696 | -46.6667 | 0.53333 | 23.9096 | -7.9096 | 7.9096 | -7.4672 | 7.4672 | strong_candidate | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 6.7 | 4.9 | TAT207 | TAT1 | -1.8 | 1.8 | 5.8 | 31.0345 | -26.8657 | 0.73134 | 6.2241 | -1.3241 | 1.3241 | -1.25 | 1.25 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 5.8 | 5.7 | TAT207 | TAT1 | -0.1 | 0.1 | 5.75 | 1.7391 | -1.7241 | 0.98276 | 5.541 | 0.15904 | 0.15904 | 0.15014 | 0.15014 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 5.5 | 4.8 | TAT207 | TAT1 | -0.7 | 0.7 | 5.15 | 13.5922 | -12.7273 | 0.87273 | 5.3133 | -0.51325 | 0.51325 | -0.48454 | 0.48454 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 5.5 | 5.8 | TAT207 | TAT1 | 0.3 | 0.3 | 5.65 | 5.3097 | 5.4545 | 1.0545 | 5.3133 | 0.48675 | 0.48675 | 0.45952 | 0.45952 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 27.4 | 19.6 | TAT207 | TAT1 | -7.8 | 7.8 | 23.5 | 33.1915 | -28.4672 | 0.71533 | 21.9361 | -2.3361 | 2.3361 | -2.2055 | 2.2055 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 6.5 | 7.1 | TAT207 | TAT1 | 0.6 | 0.6 | 6.8 | 8.8235 | 9.2308 | 1.0923 | 6.0723 | 1.0277 | 1.0277 | 0.97022 | 0.97022 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 4.2 | 3.5 | TAT207 | TAT1 | -0.7 | 0.7 | 3.85 | 18.1818 | -16.6667 | 0.83333 | 4.3265 | -0.82651 | 0.82651 | -0.78027 | 0.78027 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 7.8 | 8.7 | TAT207 | TAT1 | 0.9 | 0.9 | 8.25 | 10.9091 | 11.5385 | 1.1154 | 7.059 | 1.641 | 1.641 | 1.5492 | 1.5492 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 7.2 | 6.2 | TAT207 | TAT1 | -1 | 1 | 6.7 | 14.9254 | -13.8889 | 0.86111 | 6.6036 | -0.40361 | 0.40361 | -0.38104 | 0.38104 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 32.7 | 32.4 | TAT207 | TAT1 | -0.3 | 0.3 | 32.55 | 0.92166 | -0.91743 | 0.99083 | 25.959 | 6.441 | 6.441 | 6.0807 | 6.0807 | strong_candidate | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 7.4 | 6.3 | TAT207 | TAT1 | -1.1 | 1.1 | 6.85 | 16.0584 | -14.8649 | 0.85135 | 6.7554 | -0.45542 | 0.45542 | -0.42995 | 0.42995 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 5.4 | 4.3 | TAT207 | TAT1 | -1.1 | 1.1 | 4.85 | 22.6804 | -20.3704 | 0.7963 | 5.2373 | -0.93735 | 0.93735 | -0.88491 | 0.88491 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 5.9 | 4.9 | TAT207 | TAT1 | -1 | 1 | 5.4 | 18.5185 | -16.9492 | 0.83051 | 5.6169 | -0.71687 | 0.71687 | -0.67677 | 0.67677 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 5 | 6.3 | TAT207 | TAT1 | 1.3 | 1.3 | 5.65 | 23.0088 | 26 | 1.26 | 4.9337 | 1.3663 | 1.3663 | 1.2898 | 1.2898 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 5.2 | 5.8 | TAT207 | TAT1 | 0.6 | 0.6 | 5.5 | 10.9091 | 11.5385 | 1.1154 | 5.0855 | 0.71446 | 0.71446 | 0.67449 | 0.67449 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 4.9 | 5 | TAT207 | TAT1 | 0.1 | 0.1 | 4.95 | 2.0202 | 2.0408 | 1.0204 | 4.8578 | 0.14217 | 0.14217 | 0.13422 | 0.13422 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 8.6 | 10.5 | TAT207 | TAT1 | 1.9 | 1.9 | 9.55 | 19.8953 | 22.093 | 1.2209 | 7.6663 | 2.8337 | 2.8337 | 2.6752 | 2.6752 | mild_candidate | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 10.7 | 8.8 | TAT207 | TAT1 | -1.9 | 1.9 | 9.75 | 19.4872 | -17.757 | 0.82243 | 9.2602 | -0.46024 | 0.46024 | -0.43449 | 0.43449 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 33.9 | 31.2 | TAT207 | TAT1 | -2.7 | 2.7 | 32.55 | 8.2949 | -7.9646 | 0.92035 | 26.8699 | 4.3301 | 4.3301 | 4.0879 | 4.0879 | candidate | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 42.1 | 35.1 | TAT207 | TAT1 | -7 | 7 | 38.6 | 18.1347 | -16.6271 | 0.83373 | 33.094 | 2.006 | 2.006 | 1.8938 | 1.8938 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 16.2 | 10.1 | TAT207 | TAT1 | -6.1 | 6.1 | 13.15 | 46.3878 | -37.6543 | 0.62346 | 13.4349 | -3.3349 | 3.3349 | -3.1484 | 3.1484 | mild_candidate | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 5.2 | 5.1 | TAT207 | TAT1 | -0.1 | 0.1 | 5.15 | 1.9417 | -1.9231 | 0.98077 | 5.0855 | 0.014458 | 0.014458 | 0.013649 | 0.013649 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 3.1 | 3.8 | TAT207 | TAT1 | 0.7 | 0.7 | 3.45 | 20.2899 | 22.5806 | 1.2258 | 3.4916 | 0.30843 | 0.30843 | 0.29118 | 0.29118 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 75 | 66.9 | TAT207 | TAT1 | -8.1 | 8.1 | 70.95 | 11.4165 | -10.8 | 0.892 | 58.0663 | 8.8337 | 8.8337 | 8.3396 | 8.3396 | strong_candidate | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 8.7 | 7.9 | TAT207 | TAT1 | -0.8 | 0.8 | 8.3 | 9.6386 | -9.1954 | 0.90805 | 7.7422 | 0.15783 | 0.15783 | 0.149 | 0.149 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 8.4 | 7.8 | TAT207 | TAT1 | -0.6 | 0.6 | 8.1 | 7.4074 | -7.1429 | 0.92857 | 7.5145 | 0.28554 | 0.28554 | 0.26957 | 0.26957 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 13.2 | 10.3 | TAT207 | TAT1 | -2.9 | 2.9 | 11.75 | 24.6809 | -21.9697 | 0.7803 | 11.1578 | -0.85783 | 0.85783 | -0.80984 | 0.80984 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 7.3 | 9.5 | TAT207 | TAT1 | 2.2 | 2.2 | 8.4 | 26.1905 | 30.137 | 1.3014 | 6.6795 | 2.8205 | 2.8205 | 2.6627 | 2.6627 | mild_candidate | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 9.3 | 7.1 | TAT207 | TAT1 | -2.2 | 2.2 | 8.2 | 26.8293 | -23.6559 | 0.76344 | 8.1976 | -1.0976 | 1.0976 | -1.0362 | 1.0362 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 16 | 12.1 | TAT207 | TAT1 | -3.9 | 3.9 | 14.05 | 27.758 | -24.375 | 0.75625 | 13.2831 | -1.1831 | 1.1831 | -1.1169 | 1.1169 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 46.7 | 33.1 | TAT207 | TAT1 | -13.6 | 13.6 | 39.9 | 34.0852 | -29.1221 | 0.70878 | 36.5855 | -3.4855 | 3.4855 | -3.2906 | 3.2906 | mild_candidate | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 4 | 4.4 | TAT207 | TAT1 | 0.4 | 0.4 | 4.2 | 9.5238 | 10 | 1.1 | 4.1747 | 0.2253 | 0.2253 | 0.2127 | 0.2127 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 13 | 10.7 | TAT207 | TAT1 | -2.3 | 2.3 | 11.85 | 19.4093 | -17.6923 | 0.82308 | 11.006 | -0.30602 | 0.30602 | -0.2889 | 0.2889 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 5.1 | 4.5 | TAT207 | TAT1 | -0.6 | 0.6 | 4.8 | 12.5 | -11.7647 | 0.88235 | 5.0096 | -0.50964 | 0.50964 | -0.48113 | 0.48113 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 6.3 | 6.1 | TAT207 | TAT1 | -0.2 | 0.2 | 6.2 | 3.2258 | -3.1746 | 0.96825 | 5.9205 | 0.17952 | 0.17952 | 0.16948 | 0.16948 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 5 | 4.7 | TAT207 | TAT1 | -0.3 | 0.3 | 4.85 | 6.1856 | -6 | 0.94 | 4.9337 | -0.23373 | 0.23373 | -0.22066 | 0.22066 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 2.4 | 2.7 | TAT207 | TAT1 | 0.3 | 0.3 | 2.55 | 11.7647 | 12.5 | 1.125 | 2.9602 | -0.26024 | 0.26024 | -0.24568 | 0.24568 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 10.3 | 9.6 | TAT207 | TAT1 | -0.7 | 0.7 | 9.95 | 7.0352 | -6.7961 | 0.93204 | 8.9566 | 0.64337 | 0.64337 | 0.60738 | 0.60738 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 9.1 | 7.1 | TAT207 | TAT1 | -2 | 2 | 8.1 | 24.6914 | -21.978 | 0.78022 | 8.0458 | -0.94578 | 0.94578 | -0.89288 | 0.89288 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 11 | 10.3 | TAT207 | TAT1 | -0.7 | 0.7 | 10.65 | 6.5728 | -6.3636 | 0.93636 | 9.488 | 0.81205 | 0.81205 | 0.76662 | 0.76662 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 13.2 | 11 | TAT207 | TAT1 | -2.2 | 2.2 | 12.1 | 18.1818 | -16.6667 | 0.83333 | 11.1578 | -0.15783 | 0.15783 | -0.149 | 0.149 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 11.1 | 7.4 | TAT207 | TAT1 | -3.7 | 3.7 | 9.25 | 40 | -33.3333 | 0.66667 | 9.5639 | -2.1639 | 2.1639 | -2.0428 | 2.0428 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 5.2 | 5.3 | TAT207 | TAT1 | 0.1 | 0.1 | 5.25 | 1.9048 | 1.9231 | 1.0192 | 5.0855 | 0.21446 | 0.21446 | 0.20246 | 0.20246 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 3.9 | 4.4 | TAT207 | TAT1 | 0.5 | 0.5 | 4.15 | 12.0482 | 12.8205 | 1.1282 | 4.0988 | 0.3012 | 0.3012 | 0.28436 | 0.28436 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 123.5 | 125.8 | TAT207 | TAT1 | 2.3 | 2.3 | 124.65 | 1.8452 | 1.8623 | 1.0186 | 94.8795 | 30.9205 | 30.9205 | 29.1908 | 29.1908 | strong_candidate | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 5.4 | 4.6 | TAT207 | TAT1 | -0.8 | 0.8 | 5 | 16 | -14.8148 | 0.85185 | 5.2373 | -0.63735 | 0.63735 | -0.6017 | 0.6017 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 4.1 | 5.2 | TAT207 | TAT1 | 1.1 | 1.1 | 4.65 | 23.6559 | 26.8293 | 1.2683 | 4.2506 | 0.9494 | 0.9494 | 0.89629 | 0.89629 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 13.2 | 11 | TAT207 | TAT1 | -2.2 | 2.2 | 12.1 | 18.1818 | -16.6667 | 0.83333 | 11.1578 | -0.15783 | 0.15783 | -0.149 | 0.149 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 5.1 | 6.6 | TAT207 | TAT1 | 1.5 | 1.5 | 5.85 | 25.641 | 29.4118 | 1.2941 | 5.0096 | 1.5904 | 1.5904 | 1.5014 | 1.5014 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 6.8 | 6.3 | TAT207 | TAT1 | -0.5 | 0.5 | 6.55 | 7.6336 | -7.3529 | 0.92647 | 6.3 | 0 | 0 | 0 | 0 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 4.7 | 4.6 | TAT207 | TAT1 | -0.1 | 0.1 | 4.65 | 2.1505 | -2.1277 | 0.97872 | 4.706 | -0.10602 | 0.10602 | -0.10009 | 0.10009 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 7.7 | 10.5 | TAT207 | TAT1 | 2.8 | 2.8 | 9.1 | 30.7692 | 36.3636 | 1.3636 | 6.9831 | 3.5169 | 3.5169 | 3.3201 | 3.3201 | mild_candidate | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 4.8 | 4.1 | TAT207 | TAT1 | -0.7 | 0.7 | 4.45 | 15.7303 | -14.5833 | 0.85417 | 4.7819 | -0.68193 | 0.68193 | -0.64378 | 0.64378 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 11.5 | 9.5 | TAT207 | TAT1 | -2 | 2 | 10.5 | 19.0476 | -17.3913 | 0.82609 | 9.8675 | -0.36747 | 0.36747 | -0.34691 | 0.34691 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 4.1 | 4.9 | TAT207 | TAT1 | 0.8 | 0.8 | 4.5 | 17.7778 | 19.5122 | 1.1951 | 4.2506 | 0.6494 | 0.6494 | 0.61307 | 0.61307 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 6.1 | 5.1 | TAT207 | TAT1 | -1 | 1 | 5.6 | 17.8571 | -16.3934 | 0.83607 | 5.7687 | -0.66867 | 0.66867 | -0.63127 | 0.63127 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 9.2 | 6.3 | TAT207 | TAT1 | -2.9 | 2.9 | 7.75 | 37.4194 | -31.5217 | 0.68478 | 8.1217 | -1.8217 | 1.8217 | -1.7198 | 1.7198 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 20.9 | 13.1 | TAT207 | TAT1 | -7.8 | 7.8 | 17 | 45.8824 | -37.3206 | 0.62679 | 17.0024 | -3.9024 | 3.9024 | -3.6841 | 3.6841 | candidate | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 5.2 | 5.9 | TAT207 | TAT1 | 0.7 | 0.7 | 5.55 | 12.6126 | 13.4615 | 1.1346 | 5.0855 | 0.81446 | 0.81446 | 0.7689 | 0.7689 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 5.2 | 4.8 | TAT207 | TAT1 | -0.4 | 0.4 | 5 | 8 | -7.6923 | 0.92308 | 5.0855 | -0.28554 | 0.28554 | -0.26957 | 0.26957 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 7.8 | 8.1 | TAT207 | TAT1 | 0.3 | 0.3 | 7.95 | 3.7736 | 3.8462 | 1.0385 | 7.059 | 1.041 | 1.041 | 0.98273 | 0.98273 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 3.1 | 3.8 | TAT207 | TAT1 | 0.7 | 0.7 | 3.45 | 20.2899 | 22.5806 | 1.2258 | 3.4916 | 0.30843 | 0.30843 | 0.29118 | 0.29118 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 4.5 | 5 | TAT207 | TAT1 | 0.5 | 0.5 | 4.75 | 10.5263 | 11.1111 | 1.1111 | 4.5542 | 0.44578 | 0.44578 | 0.42085 | 0.42085 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 6.4 | 5.7 | TAT207 | TAT1 | -0.7 | 0.7 | 6.05 | 11.5702 | -10.9375 | 0.89062 | 5.9964 | -0.29639 | 0.29639 | -0.27981 | 0.27981 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 6.5 | 5.4 | TAT207 | TAT1 | -1.1 | 1.1 | 5.95 | 18.4874 | -16.9231 | 0.83077 | 6.0723 | -0.67229 | 0.67229 | -0.63468 | 0.63468 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 5.1 | 5.1 | TAT207 | TAT1 | 0 | 0 | 5.1 | 0 | 0 | 1 | 5.0096 | 0.090361 | 0.090361 | 0.085307 | 0.085307 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 9.3 | 9.8 | TAT207 | TAT1 | 0.5 | 0.5 | 9.55 | 5.2356 | 5.3763 | 1.0538 | 8.1976 | 1.6024 | 1.6024 | 1.5128 | 1.5128 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 10.8 | 10.7 | TAT207 | TAT1 | -0.1 | 0.1 | 10.75 | 0.93023 | -0.92593 | 0.99074 | 9.3361 | 1.3639 | 1.3639 | 1.2876 | 1.2876 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 9.7 | 7.1 | TAT207 | TAT1 | -2.6 | 2.6 | 8.4 | 30.9524 | -26.8041 | 0.73196 | 8.5012 | -1.4012 | 1.4012 | -1.3228 | 1.3228 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 5.8 | 5.7 | TAT207 | TAT1 | -0.1 | 0.1 | 5.75 | 1.7391 | -1.7241 | 0.98276 | 5.541 | 0.15904 | 0.15904 | 0.15014 | 0.15014 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 4.7 | 4.9 | TAT207 | TAT1 | 0.2 | 0.2 | 4.8 | 4.1667 | 4.2553 | 1.0426 | 4.706 | 0.19398 | 0.19398 | 0.18312 | 0.18312 | none | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok

## outliers
SampleID | TAT207 | TAT1 | X | Y | diff_y_minus_x | abs_diff_yx | mean_xy | rel_diff_yx_pct | bias_pct_vs_x | ratio_y_over_x | pred_y | residual | abs_residual | z_MAD | abs_z_MAD | outlier_level | outlier_basis | outlier_note | regression
| 123.5 | 125.8 | TAT207 | TAT1 | 2.3 | 2.3 | 124.65 | 1.8452 | 1.8623 | 1.0186 | 94.8795 | 30.9205 | 30.9205 | 29.1908 | 29.1908 | strong_candidate | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 75 | 66.9 | TAT207 | TAT1 | -8.1 | 8.1 | 70.95 | 11.4165 | -10.8 | 0.892 | 58.0663 | 8.8337 | 8.8337 | 8.3396 | 8.3396 | strong_candidate | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 30 | 16 | TAT207 | TAT1 | -14 | 14 | 23 | 60.8696 | -46.6667 | 0.53333 | 23.9096 | -7.9096 | 7.9096 | -7.4672 | 7.4672 | strong_candidate | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 32.7 | 32.4 | TAT207 | TAT1 | -0.3 | 0.3 | 32.55 | 0.92166 | -0.91743 | 0.99083 | 25.959 | 6.441 | 6.441 | 6.0807 | 6.0807 | strong_candidate | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 33.9 | 31.2 | TAT207 | TAT1 | -2.7 | 2.7 | 32.55 | 8.2949 | -7.9646 | 0.92035 | 26.8699 | 4.3301 | 4.3301 | 4.0879 | 4.0879 | candidate | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok
| 20.9 | 13.1 | TAT207 | TAT1 | -7.8 | 7.8 | 17 | 45.8824 | -37.3206 | 0.62679 | 17.0024 | -3.9024 | 3.9024 | -3.6841 | 3.6841 | candidate | 回帰残差 y-(slope*x+intercept) をMADで標準化 | 乖離候補であり、自動除外ではない | PassingBablok